Rough Estimate of Gas-handling Components Required: Short list

Compressor: Contacted KNF 30 Sept. ‘12

FORMERLY:  GAST MAA Diaphragm, Two-Stage	($386.90 for 1)

Mass Flow Controller: MKS πMFC  PFC-50 or PFC-60	Quote Requested	(Option2 price: $6672 for 3)

Purifier/Filter: Systems Specialties Gaskleen II Purifier	($3320 for 2)

Pump: Edwards XDS10  	($6,257 for 1)

Mixing Baffle: McMaster-Carr 12 blade Inline Pipe Mixer 	($350-$880 for 1)

Molecular Sieve: Scientific Instrument Services MT-752K Adsorption Trap	($325 for 1)

Cold-Trap:  NSCL can build one, according to J.Yurkon, but I haven’t contacted Jay Pline

FORMERLY: Scientific Instrument Services LN2 Cold-Trap	($1060 for 2)

Regulators: : McMaster-Carr Gas Pressure Control Regulator CGA 7897A16&17	($1325.80 for 5)

Pressure Transducers: MKS 622B 13TDE Baratron	($2208 for 2)

Differential Pressure Transducers: MKS 226A(21MGAGAFB1T1)	($1376 for 2)

Automated Valves: Swagelok IGC-II Pneumatic-Actuated Low-Pressure Valve

FORMERLY  Kurt J. Lesker Bellows Sealed Angle Valve (SA0100EVQF)	($2180 for 5)

Manual Valves: Swagelok IGC-II Toggle-actuated Low-Pressure Valve

FORMERLY  Kurt J. Lesker Bellows Sealed Angle Valve (SA0100MVQF)	($3360 for 12)

Storage Tank: McMaster-Carr High Pressure Inert Service HP80 (7822A14)	($303.63 for 1)

Gas Analyzer:   MKS Cirrus 2 continuous gas analyzer	($44K for 1)

Or Extrel MAX300 LG Gas Analyzer	($66K for 1)

Heat Exchanger: Necessary? If so, Exergy Sanitary 17 Series	($1800 for 1)

Ballast Tank:  McMaster-Carr ASME-Code Horizontal Pressure Tank 9426K14	($1148.45 for 1)

Total Price, minus gas analyzer & minus connecting components:			 $31,917

…with gas analyzer:									$76-98,000

Rough Estimate of Gas-Handling Components Required: Full List of Options

Compressor

Option1: Cal Supply Company Low Pressure Oil Free Rotary Claw Compressor

- Operates at 0-30psi

- Dry-running

- Requested more information (March 13) as it’s not clear this compressor can be used with recirculating gas

  - Response: What we want is too small.  Directed me to GAST

Option2: Fluitron Inc Unsure of model

- Dry-running

- Leakage (using metal seals): 1e-9 scc/s = 4e-9 standard liters per hour

- Requested more information (March 13) as only general qualities of their compressor models are quoted

  - Response: 1 stage unit can take 0.2 atm (160 torr) input at 0PSIG output for $14,000;  2-stage unit can take 0.1 atm input at 0PSIG output for $28,000.  No model # quoted

- Cost: Ana’s spreadsheet (no model quoted) says $15,300 (including metal seals)

Option3: GAST (gastmfg.com)

- Need to contact these guys to see what they think is best.  First need to find out a likely pressure drop across other components.

- Requested manual on their line of compressors:

http://www.gastmfg.com/catalogs/rotaryvane_catalog.pdf

http://www.gastmfg.com/catalogs/diaphragm_catalog.pdf

- Closest to meeting our needs: (Unsure about DAA & 15D(DC)

  - Two-Stage MAA Diaphragm:3.5bar max press, 48mbar min pres, max flow 2.4m3/hr; Noise: 69-83dB   (85dB for 8hrs a day is max allowed w/out ear plugs)

  - $386.90

  - DOA Diaphragm?  ….didn’t show-up on search by specifications we require

Mass Flow Controllers

Option1: MKS Type 1479A Metal-Sealed

- Analog control

- Full scale range anywhere from 10sccm to 20slm

- Accuracy:±1% Full Scale

- Repeatability ±0.2% Full Scale

- Leak Rate: <1e-10 scc/sec = 4e-10 standard liters per hour

- Calibrated for a SPECIFIC GAS

- “Normal Operation Pressure Differential: 10-40psig” ?? (Contacted MKS about this)

- Two year warranty

- Cost:  The Kurt J. Lesker catalog quotes the price at $1,594 + $300 for surface mount seal + $330 for the electrical connector (RS 485) $2224 each

Option2: MKS πMFC Model PFC-50 or PFC-60

- 50: “Multi-Range”

- 60: “Pressure Insensitive”

- Digital or Analog Control

- Calibrated for a specific gas, but the current gas being used can be programmed in

- Otherwise, specifications same as for 1479A

- Contacted MKS (March 13) for comparison between above 3 devices & cost estimate

  - Reminder sent april 8th.  Received reply that they’d “get back to me” on Mar 13

Purifiers

Option1: Matheson Tri-Gas PUR-Gas Cartridge Purifier System

- Max Flow rate is 3 liters per minute, so perhaps not compatible with our system

- Clean gases to 99.9999% purity

- Cartridge installed on a base-plate so as to be replaceable (without tools)

  - Check-valve closes upon cartridge replacement

  - 1,2,or 3 cartridge base plates available

- Can either get Moisture, Oxygen, Hydrocarbon, a combination of the two, or all three

- Color indicator for purity for Moisture & Oxygen purifiers

- Seems to assume inert carrier gas.  Not sure if this is for example, or the only kind of gases they recommend using this purifier system with

- See pgs. 333-336 Matheson Tri-Gas manual

Option2: Matheson Tri-Gas PUR-Gas In-Line Purifier System

- Max flow rate is 25 liters per minute

- Clean gases to 99.9999% purity

- Can either get Moisture, Oxygen, Hydrocarbon, a combination of the two, or all three

- H2O capacity of 6grams, O2 capacity of 1000grams, & Hydrocarbon capacity of 12 grams

  - Estimated life-span: >2 years

  - Higher capacity for single in-line purifiers

- Seems to assume He carrier gas.  Not sure if this is for example or really the only gas they recommend using this purifier system with

- Proper orientation is Vertical

- See pgs. 337-339 Matheson Tri-Gas manual

Option3: Entegris Inert Gas Purifier

- Meant for: N2, He, Ne, Kr, Ar, Xe, Halocarbons, Light Hydrocarbons

- Purity: <100ppb H2, CO, CO2; <100ppt H2O O2

  - Assumes <10ppm concentration of contaminants; <1ppm moisture

- Leakage: 1e-9 scc/sec = 4e-9 slm

- Must operate above 1atm, but perhaps this is ok, as we could purposely have a pressure-drop across the mass flow controller & the mixing baffle will have some pressure drop

- ~1year estimated lifetime at nominal flow rate

- Several models available for various flow ranges; pressure drops vary for each

  - 70KF: 3slm max, 1slm nominal; with 0.7psid pressure drop

  - 100KF: 20slm max, 2slm nominal; with 4psid pressure drop

  - 300KF: 50slm max, 7slm nominal; with 5psid pressure drop

    - This would be my choice. Quote requested 8 April ‘12

Option4: NuPure UltraPure PF Series Point-of-Use UHP Gas Purifier Model 600PF

- Max Flow (at 150psig): 30slm

- Liftetime: 1 year at nominal flow of 6slm

- Leak Rate: <2e-9 scc/s = 8e-9 slm for He

- Fittings: ¼”” VCR

- Used for flow gases: Ar, Ar/CH4, He, CH4, Ne, CF4, Xe, Kr, SF6,N2

- To remove impurities: H2O, O2, CO to <10ppb

- Filters contaminants above size: 0.004micron

- Requested (Mar 14) quote & pressure-drop information

Option 5: Systems Specialties Gaskleen II Purifier

- Inert Gases: GLP2INPVMM4; Flammable: GLP2SIPVMM4

- Removes H2O,O2,CO2,CO to <1ppb

- 1x109 retention of particles >0.003μm up to 5slpm

- $1660 each

Filters

Option1: Matheson Tri-Gas Model 6164-T 316 Stainless Steel High Purity Depth Filter

- Microporous fiberglass used for filtering

- 100% filtration efficiency at 0.2 micron level

- Swagelok tube fittings (VCR, VCO fittings also available (6164-V)

- Various diameters available (1/8”, ¼”, 3/8”, ½”)

- Not sure if this will work at our pressures & pressure drop.

Option2: Matheson Tri-Gas Model 6190 0.01 Micron Membrane Gas Filter

- Teflon media used for filtering

- 100% efficiency at 0.01 micron

- Swagelok tube fittings (VCR MxM fitting also available)

- Not sue if this works at our pressures & pressure drop

Option3: See Purifier Option 4

Option 4: See Purifier Option 5

Mixing Baffle (a.k.a. Static Mixer)

Option1: KoFlo Stainless steel static mixer

- $1,071 for six-blade stainless steel mixer w/ 0.37” ID, 4” length

- …overkill

Option2: McMaster-Carr Inline Pipe Mixers

- Stainless steel

- Come with 6 or 12 blades in diameters from ½” to 2” diameters

- Cost: $240-580 for 6-blade, $350-880 for 12-blade

  - Price is pipe-diameter dependent

- More blades required for lower viscosity, but perhaps our fluid is too low of viscosity for this type of mixer to be effective

Vacuum Pumps

Option1: Edwards XDS10

- Does an excellent job thus far.  I see no reason not to get another one, unless the cost ($6,257 on website) is prohibitively high

- Since shield gas (N2) won’t be recirculated, perhaps a dry pump in unnecessary.  Though I would recommend one so that there are not back-streaming concerns.

Molecular Sieve

- Option 1: Scientific Instrument Services Model & Price dependent on Pipe-diameter

  - MT-752K:

    - Traps hydrocarbons & H2O

      - NEED TO MAKE SURE CO2 is purified!

    - NW16-NW16 flange

    - 5cfm maximum flow

    - $295 for the trap + $30 for replacement zeolite charge

    - Bakeable

Cold Trap

- Option 1: Scientific Instrument Services LN2 Cold Trap

  - NW25 flanges

  - NEED TO MAKE SURE CO2 is trapped….though it should be at this temperature

  - 4” diameter, 11.5” height

  - $530.  We need two: $1060

Differential Pressure Transducers

- Option 1: MKS 226A Differential Capacitance Manometer

  - Available for several Full Scale Ranges: 0.2-1000 Torr

  - Resolution: 0.01% of Full Scale

  - Accuracy: 0.5% of Full –sacle

  - Maximum Overpressure: 120% of Full Scale

  - Maximum Line Pressure: 40psig

  - Fitting & Full-scale range depend on model #, prices vary between $605-688

  - I think we’d like the 20mbar version with bi-directional calibration

    - NW16-KF flange fitting

    - Terminal adaptor block electrical connection

    - $688 - Wee need two: $1376

Pressure Transducers

- Option 1: MKS Type 622B Baratron

  - More or less equivalent to the 622A we already have (622A no longer made)

  - Full Scale Ranges between 1 and 1000 Torr

    - Accuracy 0.25% of reading for 1-1000 Torr (Model: 622B13TDE)

    - 0.15% for 10-1000 Torr optional upon request

      - Seemingly unnecessary as our 622A does the job just fine

  - NW16KF flange

  - 622B13TDE price: $1104.  Counting the ballast tank, we need 2: $2208  (though maybe we don’t need something so accurate for the ballast tank)

Automated Valves

- Option 1: Kurt J. Lesker Company Bellows Sealed Angle Valves

  - Stainless Steel

  - High Conductance

  - Good to 10-9 torr; Leak less than 2e-9 scc/sec

  - 24V DC

  - KF16 or KF25 size

  - Assuming KF25: $436, Part number SA0100EVQF

    - We need 4: $1744

- Option 2: McMaster-Carr Sure-Seal Stainless Steel Solenoid Valves

  - NOT SURE WHAT VACUUM THEY’RE GOOD TO

  - Used for inert gases

  - For “low flow”

  - Buna-N seal

  - Connection type: NPT female

  - Pipe Size: 1/8” or ¼”; 1-5/8” x 2-7/8”

  - 120V AC or 24V DC

  - Price: $82.36-90.27 depending on details

  - Max PSI: 10-450 depending on details

  - Assuming we’ll use ¼ 24VDC 10PSI max: Product # 4639K163: $84.05

    - We need 4: $336.20

- Option 3: Swagelok Pneumatic-Actuated Low-Pressure valve with VCR fitting (free-floating)

  - For normally open, add –O to product # or –C for normally closed

  - All options have ¼” fitting:

    - Female VCR: 6LVV-DPFR4-

    - Male VCR rotatable: 6LVV-DPMR4-

    - Male VCR integral: 6LVV-DPVR4-

- Option 4: Swagelok IGC-II surface mount Pneumatic-actuated low-pressure valve

  - For normally open, add –O to product # or –C for normally closed

  - 1/8”-27 NPT inlet

  - 2-port normally opened: 6LVV-MSM-DP-2-P-O

  - 3-port normally opened: 6LVV-MSM-DP-3-P-O

  - ADD Indicator Switch:

    - 24” wire with in-line clip

    - Add M to product # for normally open, M-2 for normally closed

Manual Valves

- Option 1: Kurt J. Lesker Company Bellows Sealed Angle Valve

  - Stainless steel

  - High conductance

  - Good to 10e-9 torr; Leak less than 2e-9scc/sec

  - KF10 up to KF50 size

  - Assuming KF25: $280, Part number SA0100MVQF:   We need 13: $3640

- Option 2: Swagelok IGC-II surface-mount toggle-actuated low-pressure valve

  - 2-port: 6LVV-MSM-DPT-2-P

  - 3-port: 6LVV-MSM-DPT-3-P

Regulators

- Option 1: McMaster-Carr Gas Pressure Control Regulator CGA 580:0-100 & CGA320:0-125

  - 580: 0-100: Product Number 7897A16

    - “Standard Duty”: steel & brass w/ two 2” diameter gauges & neoprene diaphragm

    - Two-Stage: Ensures a constant, steady delivery pressure, despite inlet fluctuation (One-stage is a little more than half the price)

    - 5/8”-18 right hand female

    - For Argon, Helium, Nitrogen

    - $265.16 each.  We need two --$530.30

  - 320: 0-125: Product Number 7897A17

    - Same as above, but for Carbon Dioxide

    - $265.16 each.  We need two  $530.30

  - For all 4 required: $1060.60  ….5th required for storage tank?: $1325.80

Storage Tank

- Option 1: McMaster-Carr High Pressure Inert Service HP80 (Product No. 7822A14)

  - 2,265 liter capacity (at stp)

  - Used for Argon, Argon/CO2 blends, nitrogen, and helium

  - DOT 3AA-2015 standards: Max pressure 2,015PSI

  - $303.63

Heat Exchanger

- Option 1: Xchanger Model LC-12-1, ref data sheet #107241

  - 22”x25”x27”;  200lbs

  - …gave us quote assuming 300l/min flow instead of 5 as requested!!!!

  - Cost: $4,470

- Option 2: Exergy Sanitary 17 Series Shell & Tube

  - 0.12psi (6 torr) pressure drop

  - Cost: $1,800

- Option 3: Couldn’t we just get away with running the gas through a helical tube?

Gas Analyzer

- Option 1: Extrel MAX300 LG Gas Analyzer

  - Mass Range: 2-250amu

  - 165lbs (75kg);   60cm x 49cm x 68cm

  - NW 100CF (6 inch) flange!!!!!

  - Dual Faraday/Electron Multiplier detector

    - Upper Detection Limit: 100%

    - Lower detection limit: 10ppm Faraday, 10ppb Electro Multiplier

  - Communication via RS232

  - Software 1 (Merlin): Run in scan mode or single-ion monitoring mode

    - 80 microseconds per sample in scan mode

      - “Up to 72 ions with 20 scan segments”

    - 5ms per ion in single-ion mode

  - Software 2 (Questor 5): Analysis Rate: 400ms per component

  - Precision: ±0.0025 absolute; 30 Day Stability: ±0.005 absolute

  - Calibrated at factory with SF6 & again on-site with gases of our choosing

  - $66 minimum (see full quote)

  - Requires “very little gas”

- Option 2: MKS Cirrus 2 continuous gas analyzer

  - Continuous analysis

  - <100ppb detection of contaminants

  - 250 data points per second

  - 64.5 cm x 41 cm x 35cm;  34.5kg

  - 0.2sccm gas-consumption

  - $44,000

    - Includes travel, installation, + 1 day of training

- Option 3: SRI Instruments Multiple Gas Analyzer Gas Chromatograph

  - <100ppb detection for most contaminants

  - (Would need to be calibrated b/c a gas chromatograph)

  - $16,275

Ballast Tank

- Option 1: McMaster-Carr ASME-Code Horizontal Pressure Tank 9426K14

  - Need 340 liters worth of storage

    - P1V1=P2V2, P1=550torr  V1=500liter,  P2=800torr V2=340liter

  - ..so we go for tank with 120gallons internal volume: 454liters

  - Choose epoxy-lined steel: Model 9426K14

  - 24”x67”x31”

  - 7 openings, NPT Female type

  - Good for down to 270torr

  - Cost: $1128.45