Procedure to Set a Pressure Set-point on the MKS T3BiA Valve

* Note that most of these steps do NOT have to be done each time.  However performing the checks is a good idea.  The only essential steps are those that explicitly set the set-point and precision

Enter the following commands into PuTTy to communicate with the valve. These are also used through RS232 communication via LabView.  Bold are entered exactly, italics indicates you need to enter a custom number.  Any command with R# is a diagnostic which returns information.  Two methods of pressure control are available, model based and proportional-integral-derivative.  Commands specific to model based are orange; commands specific to PID are purple.

- Set pressure control type (see T3BiA Manual for explanation)

  - V0   (for model based)

  - V1   (for PID)

  - R51

- Set pressure units to Torr

  - F00

  - R34

- Set pressure range for high sensor (labeled on valve) to 1000Torr

  - EH10

  - R33

- Use LEARN to get system characteristics for Model Based control (while gas is flowing through the system)

  - Enter chamber volume

    - CAL1234

    - SVO588   (588 liters estimated for the tank)

    - RVO

    - USR

  - Leave volume estimator on

    - CAL1234

    - SVE1

    - RVE

    - USR

    - ROM

  - Request suggested flow

    - RLE

  - Set actual flow                                                                                                                *optimize

    - SLF120   (120 liters at STP per minute estimated)

  - Learn the valve

    - L

  - Review pump speed data

    - RCD2

- Set pump speed compensation parameters

  - SUE1

  - RUE

- Set speed-up compensation

  - SUT0.02    (0.02 second delay for baratron estimated)                                           *optimize

  - RUT

- Set speed-up filter

  - SUF0.01    (# should be 3-10 times less than number for SUT)                              *optimize

  - RUF

- Make a pressure set-point

  - T11   (1st number indicates set-point (1-6).  2nd indicates pressure control)

  - R26   (# indicates setpoint (26 for 1, 27 for 2, etc.) )

- Make this set pressure to be 100 Torr (i.e. 10% Full scale range)

  - S110   (1st number indicates set-point (1-6).  2nd indicates pressure as % full-scale range)

  - R1   (# indicates set-point (1-6))

- Set rate at which flow adjusts to desired pressure

  - I1100   (1st number indicates set-point (1-6).  2nd indicates speed as %  [leave at 100])

  - R15   (# indicates set-point (15 for 1, 16 for 2, etc.) )

- Set threshold to within 0.1% desired pressure                                                                        *optimize

  - P10.01   (1st number indicates low-threshold for process limit 1.  2nd number is % of full scale range.  To control w/in 0.1% this number is: 0.001 *(Pdesired/PFullRange)*100 ). Limited to 4 decimal places

  - R110  reports the above number

  - P20.01   (Same as above, but for high-thresh limit.  Sub 1 & 2 for 3 & 4 for 2nd set-point). Limited to 4 decimal places

  - R120 reports the above number

- Activate set-point

  - D1   (# indicates set-point (1-6))

  - R7   (# indicates set-point (7 for 1, 8 for 2, etc.) )

- Report pressure  (as % full scale range)

  - R5

- Report valve position  (as % open)

  - R6

- Set Gain

  - Mx# (x=1 for set-point 1, # is between 0-32767) [Used Gain=5 for dP=0.1%P, P≥ 80Torr]

  - Rxx reports the gain (xx=46 for set-point 1, 47 for set-point 2,  etc.)

- Set Phase

  - Xx# (x=1 for set-point 1, # is between 0-32767)[Used Phase=5 for dP=0.1%P, P≥80Torr]

  - Rxx (xx=41 for set-point 1, 42 for set-point 2, etc.)