## Slide 1: Update of Prototype AT-TPC
May 6, 2010
Weekly meeting
Daisuke Suzuki, NSCL

## Slide 2: Test room update
Completed
Construction
Clean room certification (class 1,000)
To do
Utilities
Gas bottle stand installation (10 min job)
Ethernet (Office, clean DAQ, control requested)
Clean AC supply, clean DAQ network (2~3 months)
Entrance space (work order needed)
On-going
Pump exhaust connection
Room key -> NSCL standard UB-4
Accessories
Clean room cloths
Stool, chair -> MSU surplus, local warehouse

## Slide 3: Layout plan
Cabinet 2
Desk 02
Desk 01
Rack
Gas handler
Gas bottle
Desk 04
Scale = 40”/1”
Tool box
Desk 05
12 ft
Utilities
Desk 03
45.5”
77.5”
17”
97”
12”
118.5”
21.5”
21”1/4
47”3/4

## Slide 4: Test at the test bench; setup
10 x 10 cm2
w/ 17 strips
He + CO2 (10%) 1 atm, 241Am
Cathode
α

## Slide 5: New adopter board
Mesh signal amplitude was very small.
The adopter board and the T2K was not fully connected. The anode pads floated.
Part of the PCB was cut off to improve the connection.
Vmesh = -350 V, Vcathode = - 1~2 kV; Amesh = - 60 mV (old) <-> -50 mV (new)
Measured the resistance from the adopter input to the T2K ground
Old 100 M Ω <-> New > 1 GΩ
Old
New
Adopter

## Slide 6: Negative cross talk
Run 202 He+CO2(30%) 1 atm, 241Am
Event 3
Time bin
ASIC Channel
α
Cross talk signal with negative polarity
Almost identical amplitude

## Slide 7: Measure
0.1 mm
High segmentation
-> Low Cpad/Ctot
Put a capacitor between the mesh and the ground
<-> Damage risk during discharge
Capacitive coupling among the anode pads
Amplitude ~ Cpad/Ctot
STAR TPC NIM A499, 659 (2003)
Pads are coupled by the mesh.
-> Cpad is independent on the distance between the associated pads
Wire plane
Micromegas
mesh
pad
F. Sauli, CERN/94-195, 1994
0.5 mm

## Slide 8: Result
mesh
pad
Cadd
Cadd = 0
Cadd = 2 nF
The negative pulse amplitude reduced by factor of ~3
c.f. Measured capacitance of Micromegas Ctot = 1.1~1.4 nF