# Recovering HPECR from Gate Valves Closing
If MPS tripped due to gate valves closing, please follow the HPECR recovery instructions:
- On HPECR operator page, Open Gate Valve FE_ISRC2:BGV_D0675.
- Go to Microwave Amp. in the 18GHz Microwave Transmitter section.
- On the Microwave Amp page:
  - Check on HPECR Vacuum page and ensure that the BCCG's are okay
  - Reset all interlocks / safety interlocks
  - Turn on PLC commands
  - Reset all interlocks / safety interlocks again
  - Turn on BEAM power supply ISC2_HVP_D0652
  - Reset “RF OUTPUT Set point” to ZERO
  - Turn ON RF output
  - Recover the RF output power to where it was by 100V/step.
- Go back to the HPECR operator page, turn on both source high voltage and platform high voltage HVP D0652 and HVP D0677.
- Turn back on the acceleration column puller PSEL D0678.