## Slide 1: Overview of Detector Options
Detector | Min/Max Dimensions | Energy Resolution |  | Efficiency |  |  | Additional Notes
 |  | 5.9 keV | 122 keV | 9 keV |  | 70 keV | 
U-LEGe Ultra Low Energy Germanium | 30 mm^2 x 5 mm | 140 eV | 550 eV | 99% |  | 95% | The resolution at 5.9 keV of a 50 mm2 detector at 0.1 microseconds shaping time and 100 kcps count rate is typically 300 eV (FWHM).
 | 100 mm^2 x 10 mm | 150 eV | 550 eV | >99% |  | 98% | 
BEGe Broad Energy Germanium | 2000 mm^2 x 20 mm | 350 eV | 650 eV | * | * |  | *See BEGe slide for efficiencies. We’d expect close to 100% without a window, so suspect that the default carbon composite window is decreasing this. Also available with Be window.
 | 6500 mm^2 x 30 mm | 500 eV | 750 eV |  |  |  | 
 |  |  |  | * | ~82% |  | 
LEGE Low Energy Germanium | 50 mm^2 x 5 mm | 145 eV | 500 eV | 99% |  | 95% | Ge absorbtion edge shows up in LEGe efficiency but not U-LEGe? Smallest detector has 0.025 mm thick Be window, while largest has 0.5 mm Be window. Efficiency is calculated using Geant 4 and includes effects of window. See other slide for manufacturer data.
 | 2000 mm^2 x 20 mm | 400 eV | 680 eV | 94% |  | 98% | 
Si or SiLi Detectors | Up to 5 mm thick | >= 180 eV | N/A | 100 % |  | 10% ** | **Manufacturer reports higher efficiency at 70 keV than Geant4, see other slide.
Dimensions are written as <detector area> x <thickness>
Energy resolution usually only given in data sheets at 5.9 keV and 122 keV. Specified as FWHM.
Efficiency calculated as (photons which deposit all their energy in the detector)/(photons which hit the detector) based on Geant4 simulations with the exception of BEGe, since makeup and thickness of carbon composite window isn’t known. These were compared with manufacturer defined efficiencies where they were given to check for inconsistencies.

## Slide 2: Silicon Detector Efficiency Discrepancy
I think that Mirion is reporting number of photons which interact with the detector at all divided by total particles hitting detector, whereas I’m reporting number of photons in the photo peak only divided by total number of photons hitting the detector.
Above is counts vs. energy absorbed from Geant 4 w/ 5 mm Si
Photons in the Compton continuum: ~(90/0.75)*15=1800=>18%
Photons in photo peak = 973 => 9.7%
~28% total
from: http://web-docs.gsi.de/~stoe_exp/web_programs/x_ray_absorption/index.php
From U-LEGe datasheet

## Slide 3: Is Ge Absorption Edge Important?
Some literature links:
Efficiency of Ge(Li) thin window detectors near absorption edge of germanium
The response of high-purity germanium detectors to X-rays with energy in the region of the Ge K-absorption edge
Edge is at 11.104 keV
Similar figure not given for U-LEGe. Mirion stated it wouldn’t be a concern on our phone call. Maybe the dead layer on U-LEGe is very thin?
I believe all our x-ray energies are either below the absorbtion edge, or well above, so it may be irrelevant.

## Slide 4: Detector Geometry & Efficiency
If we go with U-LEGe, it’ll be very important to get the detector within 1 cm of the source
5 mm Ge results in thickness near 1, and any increase in thickness will increase the background from gamma rays

## Slide 5: Cryostats
TODO:
Window vs windowless (window doesn’t absorb much at our energies)
Other options? Need distances to detectors)

## Slide 6: Rates
Stopped beam rates from FRIB Estimated Rates
Compute number of nuclei deposited on foil which haven’t yet decayed after X seconds
Exponential decay while foil is being moved into detector, and while in detector
 |  | PAC1 | PAC2 | Full FRIB60
60Ga | stopped beam | 9.56E+02 | 4.78E+03 | 1.50E+05
 | After 1 s in beam | 9.49E+02 | 4.75E+03 | 1.49E+05
 | 100 ms after removal | 3.53E+02 | 1.76E+03 | 5.53E+04
 | 1 s after removal | 4.76E-02 | 2.38E-01 | 7.47E+00
73Kr | Stopped beam | 4.07E+05 | 3.67E+06 | 1.14E+08
 | After 1 s in beam | 9.74E+03 | 8.79E+04 | 2.73E+06
 | 100 ms after removal | 9.72E+03 | 8.76E+04 | 2.72E+06
 | 1 s after removal | 9.49E+03 | 8.55E+04 | 2.66E+06
Decays Per Second

## Slide 7: Questions for Mirion
Information on Ge absorption edge in ULEGe
Detector to window distances. Want to confirm it’s not more than a couple mm.
Can we get a larger than 100 mm^2 ULEGe? Can we get the 100 mm^2 detector with a 5mm thick depletion region to reduce gamma background?
Confirm that ~10k counts/second is ok for all the detectors we’re considering

## Slide 8: BEGe Detector Efficiency
%solid angle covered should be 0.57%
0.47%/0.57%=82% of photons that hit the detector are seen.

## Slide 9: LEGE Efficiency

## Slide 10: 70 ke, 5 mm Ge