## Slide 1: Hardware Setup for PXCT
20230831

## Slide 2
One Ge requires:
1 BNC-BNC from iPA HV Inhibit to Controller HV Inhibit (we already have)
1 BNC-BNC from Controller HV Inhibit to HV bias supply Inhibit (we already have)
1 BNC-BNC from iPA Energy output to Oscilloscope (we already have)
1 BNC-BNC from iPA Timing output to Oscilloscope
1 BNC-SMB from iPA Energy output to Pixie-16
1 BNC-SMB from iPA Timing output to Pixie-16
One Si requires:
1 Pin-Lemo from the detector to feedthrough
1 Lemo-Lemo from feedthrough to MPR-1 input
1 BNC-BNC from MHV-4 to MPR-1 bias input
1 BNC-BNC from Pulser to MPR-1 test input
1 BNC-BNC from MPR-1 Energy output to Oscilloscope
1 BNC-BNC from MPR-1 Timing output to Oscilloscope
1 BNC-SMB from MPR-1 Energy output to Pixie-16
1 BNC-SMB from MPR-1 Timing output to Pixie-16

## Slide 3
Cables:
12 6-foot long BNC-BNC cables, $37 each
https://www.cdint.com/catalog/model/BB-316DS-6
Ten 1-foot long BNC-SMB cables: $28.34 each
https://www.cdint.com/catalog/model/BSB-174LN-1
Two 2-foot long Lemo00-Lemo00 cables, $50 each
https://www.cdint.com/catalog/model/L00L00-174LN-2
Two 4-inch long Lemo00 to BNC cables, $36 each
https://www.cdint.com/catalog/model/BL00-316DS-1
The lengths on the vendor's website range from 1 ft to 50 ft. Could you please request 4 inches long?
Adapters:
Ten BNC-BNC adapters, $5.62 each
https://www.mouser.com/ProductDetail/Amphenol-RF/031-219-RFX?qs=Y5Qy%252BBDni13toX22LO8weQ%3D%3D
Four 50-ohm BNC feedthrough terminators, $5.00 each
https://www.amazon.com/Copper-Adapter-38-5mm-Through-Terminator/dp/B07G566JC7/ref=sr_1_1_sspa?crid=3KJDDAWFFDHST&dchild=1&keywords=bnc+50+ohm+terminator&qid=1626459939&s=electronics&sprefix=BNC+50%2Celectronics%2C166&sr=1-1-spons&psc=1&smid=AD5EGQQBIXBQW&spLa=ZW5jcnlwdGVkUXVhbGlmaWVyPUEzQ1BNTEZDQzQwSFlNJmVuY3J5cHRlZElkPUEwMjgxNTAyM1RJN0tDNkwxUUdFMyZlbmNyeXB0ZWRBZElkPUEwNDk0NDc2M0NXUDhPUDRZUUcxQiZ3aWRnZXROYW1lPXNwX2F0ZiZhY3Rpb249Y2xpY2tSZWRpcmVjdCZkb05vdExvZ0NsaWNrPXRydWU=

## Slide 4
50 Ohm Coax Cable
$26
$25
$37
$28
$36
$46
$96
$35
$38
$37
harsh environments
$24

## Slide 6
About Low Loss Cable
This is Belden's 7805 "low loss" cable. It's similar to RG174, but with a solid copper core and a layer of foil added under the braid. This combination causes less attenuation, providing roughly 50% more signal as measured in decibels. 7805 also works better at higher frequencies, and is rated up to 6 GHz. The foil layer also provides for better EMF shielding versus braid alone.Belden 7805 data sheet.
Belden 8216 data sheet (an example RG174/U, for comparison)
We see this cable used in fixed-place permanent installations such as wifi and GPS antenna extensions, other RF applications, low-voltage signal sources including radiation detectors and strain gauges, high-frequency or fast-rise transient signals, and other low-vibration applications in aerospace, nuclear physics, petroleum refining and chemical processing.This cable is not recommended for use in applications where it will undergo repeated cycles of flexing or vibration -- the solid core will eventually break under those conditions. You may instead want one of these stranded-core RG174 variants:
RG174 standard has a stranded core and PVC jacket.
RG174 TPR has a stranded core and TPR jacket. The TPR jacket provides for more flexibility, most useful in bench and handheld applications.
RG174 low noise has a stranded core, PVC jacket, and a carbon layer under the braid. The carbon layer reduces flexing-induced static and provides some additional shielding. It's best at lower frequencies -- see the notes on that page.

## Slide 7
About Low Noise Cable
This is a "low noise" cable, with a carbon-impregnated conductive PVC layer added around the polyethylene dielectric; this layer helps to increase shielding, dissipate static charge, and reduce mechanically-generated electrical noise as the cable is flexed. This extra protection is useful for long cable runs, low-voltage signals, high-vibration environments, in applications including aerospace, automotive, NDT, strain gauges and accelerometers, and extremely high fidelity professional audio production.Low noise cable is generally not recommended for RF use. In our own testing, we see characteristics which differ from ordinary cable, including phase delay and impedance variations with frequency. In the field, we've seen mixed results in the GHz range; low noise cables seem to provide the best results at signal frequencies in the low MHz or lower. Ultrasonic NDT, for example, uses low-noise cable with signals as high as 100 MHz, but in transient pulse trains, where phase delay is less critical. The results you get will depend on the application, and as always you will want to do your own testing before ordering large quantities.

## Slide 8
50Ohm Insert Type BNC Feed Through Terminator

## Slide 12
Pixie-16 User Manual
Hardware Setup
Model P16F-250-16 S/N 1339
Signals are digitized by 16-bit, 250 Mega Samples per Second (MSPS) ADCs.
The Pixie-16 is designed for single exponentially decaying signals. Though the Pixie-16 can work with many different signal forms, the best performance is to be expected when sending the output from a charge-integrating preamplifier or photodetectors (e.g., photomultiplier tubes) directly to the Pixie-16 without any further shaping.
The Pixie-16 can be operated in any slot from 2 to 14 of a Pixie-16 crate. Slot 1 is used by the crate controller.
Data readout rates through the CompactPCI/PXI backplane to the host computer can be up to 109 Mbyte/s.
Full speed 32-bit, 33 MHz PCI interface to host computer, up to 109 MBytes/s readout sustained.
Pulse heights are calculated to 16-bit precision and can be binned into spectra with up to 32K channels. 216 = 64K
On-board MCA spectrum from 1K to 32K bins (software selectable), up to 4.3 billion counts per bin.
MCA spectrum
32,768 bins, 32 bit deep (4.2 billion counts/bin) for each channel

## Slide 14
The Pixie-16 module communicates with a host computer through a PCI interface. The host computer is usually either an embedded PC installed in slot 1 of the Pixie-16 crate or a remote desktop or laptop that is linked to the crate via a PCI bridge that is also installed in slot 1.

## Slide 15
One of the Pixie-16 modules must be setup as the clock master. This card must be inserted in slot 2 of the crate. Exactly one clock master must be configured or else very likely the host system will fail spectacularly.
Individual Clock Mode If only one Pixie-16 module is used in the system, or if clocks between modules do not need to be synchronized, the module(s) should be set into individual clock mode. Connect pin 7 of JP101 (the clock input) with a shunt to pin 8 (loc – IN). This will use the 50 MHz local crystal oscillator of the Pixie-16 module as the clock source.

## Slide 16
The input stages of the Rev. F Pixie-16 modules are shown in Figure 1-8. Each channel’s input termination jumper (JP1_x, x=1…16), shown on the right in red circles, is used to connect the analog input signal to a 50 Ω resistor then to ground. It is installed by default at the factory, and is recommended for being used to avoid input signal reflection when using long coaxial cables to connect detector outputs to the Pixie-16 front panel connectors. However, if the detector outputs have 50 Ω output impedance, the signal will be effectively halved at the input of the Pixie-16 when the input termination jumper is installed.
If Jumper is not installed: If Relay Switch is closed, attenuation is 1:1 (1k Ω termination), and overall effective gain is 2.5; If Relay Switch is opened, attenuation is 1:4 (4k Ω termination), and overall effective gain is 0.625;
If Jumper is installed and detector preamplifier has 50 Ω output impedance: If Relay Switch is closed, attenuation is 1:1 (47.6 Ω termination), and overall effective gain is 1.25; If Relay Switch is opened, attenuation is 1:4 (49.4 Ω termination), and overall effective gain is 0.3125.
Our Pixie-16 is Rev F
I've removed the jumpers of Ch12,13,14,15 for a larger gain.

## Slide 17
Front Panel LEDs for the Pixie-16 Modules
Start Run in nscope or begin in ReadoutShell: Green LED on.
Have to reboot the crate

## Slide 22
iPA Gain×10
Analog Gain 4.0
Energy filter TPeaking 8 μs
Energy filter TGap 1.5 μs

## Slide 23
σ = 6.6 ch 169 eV
μ = 230.1 ch 5.9 keV
FWHM = 397 eV
~80 Kα X rays / s
GADGET I

## Slide 24
122
244
344
778
6

## Slide 25
Turbopump cooling fan on
Turbopump cooling fan off

## Slide 26
Turbopump cooling fan on
Turbopump cooling fan off

## Slide 27
Turbopump cooling fan on
Turbopump cooling fan off

## Slide 28
iPA control panel open
iPA control panel closed

## Slide 29
iPA control panel open
iPA control panel closed

## Slide 30
iPA control panel open
iPA control panel closed

## Slide 31
Turbopump cooling fan off
Empty
Turbopump cooling fan on

## Slide 32
iPA Control Panel status: Closed
CP5 Control Panel doesn't affect the noise level
iPA Control Panel status: Open

## Slide 33
Adding grounding cables between the detector and the NIM crate

## Slide 34
RG58
RG58 Low Noise
RG62

## Slide 35
RG62
No Al foil
RG62
With Al foil
With sleeve
RG62
With Al foil