## Slide 1
Specifications | LEGe X-ray | North/Window Ge | South/Inside Ge | ΔE Si | E Si
Detector Model | GL0510 | GX10020 | GX10020 | MSD012 | MSD026
Detector S/N | 13725 | 5593 | 5596 | 3506-3-1 | 3498-25-2
Preamplifier Model | iPA P10 | iPA SlimLine10 | iPA SlimLine10 | MPR-1 | MPR-1
Preamplifier S/N | 13000040 | 201216-0207 | 210610-0560 | 1021511 | 1021512
Cryostat Model | Cryo-Pulse 5 Plus FM | Cryo-Pulse 5 Plus SL | Cryo-Pulse 5 Plus SL |  | 
Cryocooler S/N | 2751 | 2749 | 2750 |  | 
Controller S/N | 2402 | 2358 | 2355 |  | 
Endcap Diameter | 38.1 mm | 101.6 mm | 101.6 mm |  | 
Endcap Length | 203.2 mm | 171.45 mm | 171.45 mm |  | 
Crystal Diameter | 25.0 mm | 84.8 mm | 79.8 mm | 12 mm | 26 mm
Crystal Length | 10.5 mm | 65.2 mm | 80.0 mm | 12 μm | 1 mm
Distance from Window | 5.6 mm | 6.8 mm | 6.3 mm |  | 
Window Thickness | 0.13 mm Be | 0.6 mm C | 0.6 mm C | 50 nm/300 nm | 50 nm/1000 nm
Depletion Voltage | –600 V | +4000 V | +2200 V | –1.5 V | –90 V
Recommended Bias Voltage | –1100 V | +4500 V | +4500 V | –3.0 V | –130 V
Reported Test Point Voltage | –1.70 V | –2.40 V | –0.66 V |  | 
Measured Test Point Voltage | –1.71 V | –0.40 V | –0.66 V |  | 
Bias Supply | ORTEC 660 Ch A | ORTEC 660 Ch A | ORTEC 660 Ch B | MHV-4 Ch 0 | MHV-4 Ch 1
Bias Supply S/N | 22004180 | 22004181 | 22004181 | 0721242 | 0721242
Power Supply | MNV-4 Ch 4 | MNV-4 Ch 1 | MNV-4 Ch 2 | MNV-4 Ch 4 | MNV-4 Ch 3
Power Supply S/N | 1021266 | 1021266 | 1021266 | 0122278 | 0122278
Notes: PXCT ALL DETECTORS

## Slide 2
Specifications | LEGe X-ray | North/Window Ge | South/Inside Ge | ΔE Si | E Si
Measured Leakage Current | 0-100 pA | 0-20 pA | 0-20 pA | <1 nA | 50-100 nA
Measured Leakage Current w/ Source | 10-140 pA | 50-70 pA | 50-70 pA | <1 nA | 50-100 nA
Preamp Output Polarity | Positive | Negative | Negative | Positive | Positive
Preamp Output Gain | x1, x2, x5, x10 | x1, x2, x5, x10 | x1, x2, x5, x10 | 40/200 MeV | 40/200 MeV
Preamp Output Rise Time | ~150 ns | ~250 ns | ~250 ns | ~500 ns | ~70 ns
Preamp Output Fall Time | ~120 μs | ~120 μs | ~120 μs | ~280 μs | ~280 μs
Digital Shaping Time Rise Time | 7.2 μs | 10.2 μs | 10.8 μs |  | 
Digital Shaping Time Flat Top | 0.8 μs | 1.8 μs | 1.2 μs |  | 
Capacitance at Recommended Bias | 6 pF | 23 pF | 28 pF |  | 
Maximum Controller Power | 180 W | 180 W | 180 W |  | 
Operating Controller Power | 54-70 W | 120-160 W | 90-120 W |  | 
Nominal PRTD1 | –182.6℃ | –163.6°C | –170.9°C |  | 
PRTD1 HV-inhibit Threshold | –172.6°C | –153.5°C | –160.9°C |  | 
Measured PRTD1 | –182.7℃ | –157.5°C | –168.3°C |  | 
Nominal PRTD2 | –190.7℃ | –198.3°C | –197.6°C |  | 
PRTD2 HV-inhibit Threshold | –180.7℃* | –188.2°C | –187.3°C |  | 
Measured PRTD2 | –190.7℃ | –198.5°C | –197.6°C |  | 
Cold-tip Setpoint | –185°C | –185°C | –185°C |  | 
Measured Cold-tip Temperature | –185°C | –185°C | –185°C |  | 
Recommended Cool-down Time | 12 hours | 24 hours | 24 hours |  | 
Typical Warm-up Time | 12–24 hours | 12–24 hours | 12–24 hours |  | 
Digitizer | XIA Pixie-16 Ch 1 | XIA Pixie-16 Ch 3 | XIA Pixie-16 Ch 5 | XIA Pixie-16 Ch 6 | XIA Pixie-16 Ch 8
Notes: PXCT ALL DETECTORS

## Slide 3
Digitizer Pixie-16 | S/N 1339
XIA Crate WIENER UEP6021 | S/N 0479032
 | P/N 0P00.2291U
NIM Crate N8360Y | S/N 224
 | P/N WN8360YAAAAA
Agilent IDP-15 dry scroll pump, with inlet valve, X3815-64010 | MY2120SA07
Agilent TwisTorr 84 FS Turbo pump, X3502-64001 | IT21476067
Agilent TwisTorr 84 FS-AG rack controller, X3508-64001 | IT2143C133
PCM301 Busy Bee Vacuum Gauge | S/N PCM301TUTBCA
PCM301 Busy Bee Vacuum Gauge | P/N 23B02324
WGM701 Wasp Vacuum Gauge | S/N WGM701SHFA P/N 23C01643 S/N WG7SF-100 P/N 540546539
AGC302 Wasp Gauge Controller | 23F04033