## Slide 1: Digital Data Acquisition System (DDAS) Instructions
Lijie Sun
3/12/2024

## Slide 2: Hardware Setup
Our group owns two DAQ systems: XIA crate + Pixie-16 digitizer + spdaq computer.
Insert the Pixie-16 card into slot 2 of the crate.
  Pixie-16 S/N1148 is used in Rm1039; Pixie-16 S/N1339 is used in Rm1035.
For a single digitizer configuration, connect the jumpers according to the individual clock mode (next slide).

## Slide 4
Front Panel LEDs for the Pixie-16 Modules
Start Run in qtscope or Begin Run in ReadoutShell: Green LED on.
Red LED on. You have to reboot the crate

## Slide 5: Software Setup
Rm1039 protondetector@spdaq55
Rm1035 pxct@spdaq61
The same password is used for both DAQ systems and for both local login and remote SSH access. If you have access to Rm1035 or Rm1039, you'll be able to see it.

## Slide 6: 1) QtScope
cd ~
./goqtscope
Click Boot System
It reads in a cfgPixie16.txt in ~/readout/crate_1/
1 # crateID
1 # number of modules
2 # slot for mod 0
~/readout/crate_1/Test.set

## Slide 7
You should see these messages after clicking Boot
Rm1039
Rm1035

## Slide 8
Please refer to the following slides for "correct" configurations.

## Slide 9
Click Channel DSP and configure filter parameters, and then click "save settings" as a set file in ~/readout/crate_1
Please refer to the following slides for "correct" configurations.

## Slide 13: 2) ReadoutShell
cd ~
./goreadout
Click Start
Click Begin
Click End

## Slide 16
Assigned by crate_1/ReadoutCallouts.tcl
Hard drive full
increase when collecting data.

## Slide 17
Numbers increase when collecting data.

## Slide 19: 3) DDASdumper
cd ~
./goddasdumper runnumber
goevt
  You'll see the root files.

## Slide 20
dchan->Draw("energy>>(65536,0,65536)","chanid==13")